IMPACTS 2023 Cloud Probes Status and Updates
David Delene, Jennifer Moore, Christian Nairy, and Michael Willette

---
IMPACTS Cloud Probes
King
RICE
Hawkeye
HVPS3-A
HVPS3-B
Hotwire Boom
WCM
Empty
or
RICE-B
PHIPS
or
CAPS
CDP
2DS
King Probe

---
IMPACTS 2022 Cloud Probes
Measurement Size Range
Measurement
Method
1 um	10 um	100 um	1 mm	10 mm
 CDP (2.0-50.0 μm)
Forward Scattering Probes
FCDP (2.0-50.0 μm)
CAPS/CAS (0.5-50.0 μm)
CAPS/CIP (25 um to 1.55 mm)
CPI (2.3 um to 1 mm)
Imaging
Probes
PHIPS (3 um to 2 mm)
2DS (10/50 um to 1.280/6.400 mm)
(2) HVPS-3 (150 um to 19.2 mm)
King (wire) (LWC for 5 to 200 um)
Hot-wire Probes
WCM (wire) (LWC for 5 to 200 um)
Vibrating Cylinder
(2) Rosemount Ice Detector (RICE) (SLW)

---
Instrument Status
Quality assurance is being completed for 2022.
CDP and 2D-S (including leak) have been repaired.
All probes have ben checked and calibrated.
WCM and CDP have been tested in the University of North Dakota lab.
The review of the 2D-S and HVPS3 is pending.
New UPS has been installed in cloud probe instrument rack.

---
2022 Deployment Science Flights
	1/14	1/17	1/19	1/28	1/29	2/3	2/4	2/8	2/13	2/17	2/19	2/25
CDP	Failed	NA	NA				PA
WCM							PA	NA
King							PA
RICE							PA
Hawkeye
2D-S	NA	NA	NA	Failed	NA	NA	NA
HVPS-3A		NA	NA
HVPS-3B
Notes:
NA - Not Available
PA - Partially Available
1/14 - CDP Failed Upon Take-off
1/29 - 2D-S Failed
2/4 - M300 Data System Failed Mid-flight
2/8 – 2/25 – 2D-S Leak Issue

---
2023 Upload
Reasonable to plan for three weeks as all past uploads required this much time.
We required sufficient time to have any issues resolved and possible repair completed before December test flights.
University of North Dakota personnel are available October 3 – 21.
Populated racks and all probes will be sent to arrival 1 week before upload.