Atmosphere Sciences
Instrumentation Lab
Organized by David Delene - November 20, 2008
Overview of Instruments – 3:00 Clifford Hall 470
Lab Demonstrations – 4:00 - 6:00 Clifford Hall 423
Currently there are many state of the art scientific instruments set up in the Atmospheric Sciences Instrumentation Lab (CH 423). Many students are currently working with these instruments, either as part of the Measurement System course (AtSc 535) or as student researchers.

---
Instruments and Presenters
Cloud Particle Imager (CPI) – Andrea Neumann
Passive Cavity Aerosol Spectrometer Probe (PCASP) – Kelsey Watkins
King Hot Wire Liquid Water Probe – Robert Mitchell
High Volume Precipitation Spectrometer (HVPS) – Matt Ham
Dew Point Temperature Sensor – Angelle van Oploo
Pressure Transducer – David Keith
Mass Flow Controller – Tim Logan
Rain Gauge – Dan Adriaansen
UWyo Cloud Condensation Nuclei Counter (CCNC) – Gokhan Sever
Condensation Particle Counter (CPC) – Zhe Feng
Aircraft Data System (M300) – Randall Johnson
Forward Scattering Spectrometer Probe (FSSP) – Christopher Kruse

---
Cloud Particle Imager (CPI)
Andrea Neumann
The CPI uses two lasers to detect a particle and then uses a 3rd laser to take a picture.
Each picture is 1024 x 1024 pixels with each pixel representing 2.3 microns.

---
Cloud Particle Imager (CPI)
Example Pictures
Pictures courtesy of:
 Atmospheric Radiation Measurement program: http://www.arm.gov/images/isdac_cloudparticle.jpg
SPEC Inc. http://www.specinc.com/images/cpi_images.jpg

---
Passive Cavity Aerosol
Spectrometer Probe (PCASP)
Kelsey Watkins
Makes aerosol measurements of particles from 0.10 - 3.00 μm, with a resolution in the smallest channels of 0.02 μm
Aerosol size spectrum is measured at 1 Hz frequency, using a combination reflecting-refracting imaging system
Worked with the PCASP, doing calibration checks and data analysis.

---
King Hot Wire Liquid Water Probe
Robert Mitchell
Heated wire in free air stream (125°C, 165 [King])
Conductivity of air removes heat in dry conditions.
Additional heat is removed due to vaporization of liquid water when in cloud.
Difference in energy is computed into a liquid water content.
Accurate from .1g/m3 – 5g/m3
Issues:
Large drops splash off prior to vaporization
Ice drops may bounce off or not vaporize completely
Very small drops deflected around probe due to slipstream.

---
King Hot Wire Liquid Water Probe
Robert Mitchell
Challenge:
Power varies with Temperature, Pressure, Airspeed, and Others?
Very sensitive to Grounding issues and “Noise”

---
High Volume Precipitation Spectrometer (HVPS)
Matt Ham
Measures size, shape and concentration of precipitation sized particles
Particles create shadows which are converted to binary ones
Electronically compatible with PMS 2D-P
Samples at rate proportional to aircraft speed
4.5x20 cm opening

---
High Volume Precipitation Spectrometer (HVPS)
Sample Output

---
Dew Point Temperature Sensor
Angelle van Oploo
External Sampling Probe
Optical Chilled Mirror primary measurement of dew point temperature.
Control Unit - Data Collection and Automatic Balance Cycle
Down Points - Frost/Dew Temperature

---
Pressure Transducer
David Keith
Pressure sensing diaphragm and two coils oriented normal to the diaphragm’s axis determine output voltage used to calculate voltage.
Validyne P55 differential pressure gauge has output range of +/- 5v DC.
Aids in determining 3 dimensional wind field.

---
Mass Flow Controller
Tim Logan
Mass Flow Controller regulates the flow of air to maintain a set flow rate.
Calibration of mass flow controllers can be performed using a bubble flow meter.

---
Tipping Bucket Rain Gauge
Dan Adriaansen
Measures rainfall and rain rate in 0.01” (0.25mm) increments
8” diameter funnel provides large sampling area while still allowing small tipping mechanism
Funnel helps prevent evaporation inside instrument
Each tip creates a 100ms switch closure
Closure can be recorded as ‘tip’
Accuracy of 0.5% at 0.5”/hr

---
Tipping Bucket Rain Gauge

---
UWyo Cloud Condensation
Nuclei Counter (CCNC)
Gökhan Sever
CCN; particles capable of forming cloud droplets under supersaturation conditions.
Measures droplet formation ability of aerosol particles.
Supersaturation is created by means of temperature gradient between two plates.
Supersaturation range is ~0.2% - 2.0%.
Supersaturation values are adjusted step-by-step.
Activated particles are counted as a function of SS
Activated droplets are counted in chamber by measuring the scattered light.
Analysis of data from S. Arabia and Mali Rainfall Enhancement projects.

---
Condensation Particle Counter
Zhe Feng
Air sample drawn continuously, mixed with alcohol vapor (heated saturator)
Cooled condenser makes supersaturation → alcohol condense on particles
Condensational growth makes larger droplets, pass through optical detector
Laser source and photodetector measures scattered light from particles
Detects airborne particles
D>10 nm at 1.0 L/min flow rate, concentration up to 104 cm-3
Flow rate controlled using an internal critical orifice
Going to work with it in Measurement System lab.
Alcohol Bottle
Inlet

---
Aircraft Data System (M300)
Randall Johnson
Data Acquisition System to gather data from instruments
Based on QNX operating system
Real-time computation and display
Interfacing with M300 is easy
Support for: analog, digital, 1D probes, 2D probes, GPS … etc.
Versatile: user configurable probe channel tables
PCASP calibration check

---
Forward Scattering
Spectrometer Probe (FSSP)
Christopher Kruse
Droplets are sized by measuring the amount of light scattered during interaction through a focused laser beam.
Nominal range for cloud droplet sampling is 2 – 47 µm.
Measures the cloud droplet spectrum (typically in 15 or 20 bins) in developing liquid water clouds.

---
Question
Cloud Particle Imager (CPI) – Andrea Neumann
Passive Cavity Aerosol Spectrometer Probe (PCASP) – Kelsey Watkins
King Hot Wire Liquid Water Probe – Robert Mitchell
High Volume Precipitation Spectrometer (HVPS) – Matt Ham
Dew Point Temperature Sensor – Angelle Van Oploo
Pressure Transducer – David Keith
Mass Flow Controller – Tim Logan
Rain Gauge – Dan Adriaansen
Uwyo Cloud Condensation Nuclei Counter (CCNC) – Gokhan Sever
Condensation Particle Counter (CPC) – Zhe Feng
Aircraft Data System (M300) – Randall Johnson
Forward Scattering Spectrometer Probe (FSSP) – Christopher Kruse
Lab Demonstrations – 4:00 – 6:00 Clifford Hall 423